## SYLLUBUS
1 | INTRODUCTION |  |  |  |  | REFERENCE
| OUR PRODUCTS - ANALYZER(LSE V3,LSE S V3,ECO SV,ECOD,ECOV,ECO-C,ECO-CP,VIBRO AND EXTERNAL DISPLAY) |  |  |  |  | http://ecoursesonline.iasri.res.in/course/view.php?id=347
| MILK TYPES AND PARAMETERS |  |  |  |  | https://www.dairyknowledge.in/content/types-milk
MILK TESTING METHOD-GERBER FAT,CLR-LACTOMETER
| ANALYZER TESTING TECHNOLOGY | FUNCTION OPERATION AND TESTING METHOD
2 | MACHINES
2.A | LSE V3 | ANALYZER - HARDWARE OPERATIONS | LCD AND CONNECTION DETAILS | OPERATIONS AND TROUBLESHOOTING
|  |  | KEY PAD AND CONNECTION DETAILS | OPERATIONS AND TROUBLESHOOTING
|  |  | PUMP AND OPERATION DETAILS | MOTOR FIXING DETAILS AND ROTATION RELATED POINTS | OPERATIONS AND TROUBLESHOOTING
BELLOW PCB FIXING DETAILS AND TROUBLESHOOTING
BELLOW OPERATIONS AND IMPORTANT POINTS
|  |  | RS232 AND CONNECTION DETAILS | OPERATIONS AND TROUBLESHOOTING
|  |  | MAIN BOARD | POWER SUPPLY SIDE | OPERATIONS AND TROUBLESHOOTING
|  |  |  | CONTROLLER SIDE OPERATIONS | OPERATIONS AND TROUBLESHOOTING
|  |  |  | MOTOR DRIVER | OPERATIONS AND TROUBLESHOOTING
|  |  |  | LCD DRIVER | OPERATIONS AND TROUBLESHOOTING
|  |  |  | HEATER DRIVER | OPERATIONS AND TROUBLESHOOTING
|  |  |  | BUZZER DRIVER | OPERATIONS AND TROUBLESHOOTING
|  |  |  | ANALYSE SIGNAL TRANSMITTER SIDE | OPERATIONS AND TROUBLESHOOTING
|  |  |  | ANALYSE SIGNAL RECIEVER SIDE | OPERATIONS AND TROUBLESHOOTING
|  |  |  | SIGNAL PROCESSING SIDE - OP-AMP | OPERATIONS AND TROUBLESHOOTING
|  |  |  | NEGATIVE VOLTAGE IC DRIVE | OPERATIONS AND TROUBLESHOOTING
IMPORTANT TEST POINTS AND VOLTAGE RANGES DETAILS
|  |  | SENSOR BOARD | COILS CONNECTIONS AND OPERATIONS | OPERATIONS AND TROUBLESHOOTING
|  |  |  | COIL SUPPLY REGULATOR OPERATIONS | OPERATIONS AND TROUBLESHOOTING
|  |  |  | SIGNAL SENSING METHOD AND TROUBLESHOOTING | OPERATIONS AND TROUBLESHOOTING
|  |  |  | BOX TEMPERATURE AND CIRCUIT OPERATIONS | OPERATIONS AND TROUBLESHOOTING
SENSOR CABIN FIXTURE DETAILS AND IMPORTANT POINTS
SILICON TUBE FIXING DETAILS AND IMPORTANT POINTS
|  |  | PRE HEATER | CONNECTION DETAILS AND IMPORTANT POINTS | OPERATIONS AND TROUBLESHOOTING
|  | SOFTWARE DETAILS | SOFTWARE FINAL VERSIONS AND CHANGES
|  |  | TESTING MILK TYPES ( CHANNELS ) | CHANNEL SEPARATIONS REASON AND IMPORTANT POINTS
|  |  | AUTO CHANNEL OPERATIONS DETAILS | ENABLE/DISABLE DETAILS
|  |  | MILK SAMPLE TEST DETAILS (OPERATIONAL) | RELATION WITH SAMPLE MILK TESTING METHOD AND READING PARAMETERS.
|  |  | CHANNELS CALIBRATION | IMPORTANTS OF THE CHANNEL CALIBRATION AND OPERATIONS
IMPORTANT POINTS OF THE CALIBRATION PROCESS
MANUAL CORRECTION METHOD AND AUTO CALIBRATION DETAILS
|  |  | TEMP.COMPANSATION | TEMP.COMPANSATION IMPORTANTS AND OPERATIONAL DETAILS
CHANNEL OFFSET CORRECTION AND RESOLUTION SET DETAILS
|  |  | WATER ZERO SET OPERATION DETAILS | HOW TO SET THE WATER ZERO AND IMPORTANT NOTED POINTS
CLR EQUATION DESCRIPTION DETAILS
CHANNEL ENABLE/DISABLE OPERATION (PROFILE)
AUTO CLEANING CYCLE TEST OPERATION
SENSOR QUALITY TEST (SIGNAL COUNT,READING TEMPERATURE)
COLLECTION METHOD-NORMAL,REMOTE
|  |  | COLLECTION STORAGE AND SEND TO PC DETAILS | SEND DATA FORMAT SETTING DETAILS
|  |  | RS232 SERIAL PORT FORMAT SETTING DETAILS | PURPOSE OF THE DIFFERENT OUT PUT FORMATS DETAILS
DISPLAY SHOWING MILK PARAMETERS SETTING DETAILS
THERMAL PRINTER CONNECTIVITY SETTING DETAILS
SAMPLE TEST TIME SHOWING SELECTION ( TIME BAR/TIME COUNT)
DAILY/WEEKLY CLEANING DETAILS
2.B | LSE S V3 , ECO SV AND ECO V | ANALYZER HARDWARE AND SOFTWARE FUNCTIONS | SAME
|  | ADDITIONAL HARDWARE | SOLAR CHARGER | FUNCTIONAL OPERATION OF CHARGER
WIRING DETAILS OF THE CHARGER
CHARGER SPECIFICATIONS-CHARGING VOLTAGE,CURRENT AND CUTOFF RANGE
VOLTAGE DRIVER SIDE DETAILS
|  |  | STIRRER BOARD | FUNCTIONAL OPERATION OF STIRRER
SIGNAL GENERATOR CIRCUIT DETAILS
STIRRER GUN OPERATION DETAILS
GUN DRIVING CIRCUITARY DETAILS(TRANSFORMER)
2.C | ECOD | ANALYZER BOARD OPERATIONS | SAME
|  | DPST BOARD DETAILS - HARDWARE | FUNCTIONAL OPERATIONS OF DPST BOARD
CONTROLLER OPERATIONS
|  |  | EXTERNAL INTERFACING PERIPHERALS DETAILS | WEIGHNING SCALE,EXTERNAL DISPLAY,COMPUTER,MULTIMEDIA KEYBOARD AND PENDRIVE
|  |  | INTERNAL PERIPHERAL DETAILS | LCD,MATRIX KEYBOARD,TIMER RTC,THERMAL PRINTER
|  | SOFTWARE SIDE | SD CARD FORMATING - USE MATRIX KEYPAD
|  |  | ANALYZER BOARD INTERFACE SKIPPING DETAILS | DUE TO ANALYZER BOARD FAILURE
|  |  | MATRIX KEYBOARD SHORTCUT DETAILS | SAMPLE SELECTION AND TESTING
INTERNAL PERIPHERAL STATUS WINDOW SELECTION DETAILS
COLLECTION SUMMARY PRINT
SAMPLE TEST RESULT SECOND PAGE DISPLAY SELECTION DETAILS
MANUAL ENTRING SHORTCUT SELECTION - RATE,WEIGHT AND ANALYZER PARAMETERS
|  |  | MATRIX KEY BOARD FUNCTIONAL OPERATIONS | MENU FUNCTIONS,TARE OPERATION,FARMER ADDING,SALES ENTERING,MANUAL CLEANING,PRINTOUT AND VIBRO CONTROL
|  |  | RATE CHART RELATED DETAILS | CHART TYPES | COW,BUFFALO AND MIX
|  |  |  | AUTO CHART | ENABLE/DISABLE AND AUTO FAT LIMIT SETTINGS
TEST RESULT RATE SELECTION SETTINGS DETAILS
|  |  |  | CHART GENERATION | 5-METHODS GENERATION DETAILS
|  |  | WEIGNING SCALE SELECTION | MANUAL SETTINGS OF THE SCALE DATA SELECTION | FROM THE SCALE SERIAL PORT DATA WITH AND WITHOUT DATA SELECTION SETTINGS
|  |  | WEIGHNING SCALW TARE SETTINGS | FUNCTIONAL OPERATION | BEFORE AND AFTER TARE SETTINGS

## CLR_EQUATION
Sl. No. | State | Formula | Remarks
1 | Gujarat | CLR/4 + 0.21(Fat%) + 0.36 | Lactometer calibrated at 21 deg C
2 | West Bengal | CLR/4 + 0.21(Fat%) + 0.36 | Lactometer calibrated at 15.5 deg C
3 | Orissa | CLR/4 + 0.21(Fat%) + 0.36 | Lactometer calibrated at 15.5 deg C
4 | WAMUL, Assam | CLR/4 + 0.20(Fat%) + 0.66 | Lactometer calibrated at 29 deg C
5 | Maharashtra | CLR/4 + 0.21(Fat%) + 0.36 | Lactometer calibrated at 21 deg C
6 | Goa | CLR/4 + 0.21(Fat%) + 0.36 | Lactometer calibrated at 21 deg C
7 | MP | CLR/4 + 0.20(Fat%) + 0.70 | Lactometer calibrated at 29 deg C
8 | UP | CLR/4 + 0.20(Fat%) + 0.29 | Lactometer calibrated at 15.5 deg C At DCS level it is CLR/4+0.20(Fat%) + 0.5 (Lactometer at 29 deg C)
9 | Punjab | CLR/4 + 0.20(Fat%) + 0.29 | Lactometer calibrated at 15.5 deg C
10 | Haryana | CLR/4 + 0.20(Fat%) + 0.29 | Lactometer calibrated at 15.5 deg C
11 | BIS | CLR/4 + 0.25(Fat%) + 0.44 | Lactometer calibrated at 27 deg C
12 | Tamil Nadu | CLR/4 + 0.20(Fat%) + 0.36 | Lactometer calibrated at 21 deg C
13 | Karnataka | CLR/4 + 0.25(Fat%) + 0.35 | Lactometer calibrated at 27 deg C
14 | A P | CLR/4+.21(Fat%)+.36 | Lactometer calibrated at 21 deg C
15 | Kerala | CLR/4 + 0.20(Fat%) + 0.50 | Lactometer calibrated at 29 deg C
16 | Pataliputra(Bih) | CLR/4 + 0.20(Fat%) + 0.14 | Lactometer calibrated at 15.5 deg C
17 | Rest of Bihar | CLR/4 + 0.20(Fat%) + 0.66 | Lactometer calibrated at 29 deg C

## TECHNICAL CHECK
SI.NO | PSC: | PRODUCT/PARTS | COMPLAINT | CHECK-1 | ACTION | CHECK-2 | ACTION | CHECK-3 | ACTION | CHECK-4 | ACTION | CHECK-5 | ACTION | CHECK-6 | ACTION | CHECK-7 | ACTION | CHECK-8 | ACTION | CHECK-9 | ACTION | CHECK-10 | ACTION | CHECK-11 | ACTION | CHECK-12 | ACTION
|  | VIBRO | NOT WORKING (NOT ON LED) | CHECK THE FUSE | REPLACE FUSE | CHECK THE POWER SUPPLY | REPLACE POWER SUPPLY | CHECK THE SUPPLY VOLTAGE (11v TO 14V) | REPLACE POWER SUPPLY | CHECK THE GUN ELEMENT SHORT | REPLACE THE GUN | REPLACE THE DC CONNECTOR | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  |  | NOT WORKING (LED JUST ON AND OFF) | CHECK THE GUN ELEMENT SHORT | REPLACE THE GUN | CHECK DC CONNECTOR | REPLACE THE DC CONNECTOR | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  |  | MACHINE ON BUT THE VIBRATION NOT PRESENT | CHECK THE SUPPLY VOLTAGE (11v TO 14V) | REPLACE POWER SUPPLY | CHECK THE GUN ELEMENT SHORT | REPLACE THE GUN | CHECK THE FREQ.POT AND CABLE | REPLACE THE FREQ.POT | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  |  | MACHINE CONTINIOUSLY VIBRATING | CHECK THE TIMER POT. AND CABLE | REPLACE THE TIMER POT. | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  |  | LOW VIBRATION | CHECK THE SUPPLY VOLTAGE (11v TO 14V) | REPLACE POWER SUPPLY | CHECK THE FREQ.POT AND CABLE | REPLACE THE FREQ.POT | CHECK THE STIRRER BOARD FREQ.TRIMPOT ADJUST | ADJUST THE FREQ.TRIMPOT | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  |  | VIBRATION INDICATION LED NOT WORKING | CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY | REPLACE THE LED AND CABLE ASSY. | CHECK THE VOLTAGE OF THE INDICATION LINE | REPLACE THE STIRRER BOARD
|  |  | POWER LED NOT WORKING | CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY | REPLACE THE LED AND CABLE ASSY. | CHECK THE VOLTAGE OF THE INDICATION LINE | REPLACE THE STIRRER BOARD
|  |  | STIRRER WORKING ONLY THE ADAPTER PLUGGING TIME.(VOBRO NOT WORKING BY THE PUSH SWITCH) | CHECK THE KEY FUNCTIONALITY AND CABLE ASSY. QUALITY. | REPLACE THE KEY | CHECK THE STIRRER BOARD | REPLACE THE STIRRER BOARD
|  | SOLAR CHARGER BOARD | NOT PRESENT THE BATTERY CHARGING VOLTAGE - 0V | CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V | REPLACED THE CHARGER ADAPTER | CHECK THE CHARGER FUSE | REPLACE THE FUSE | CHECK THE CHARGER ADAPTER TO SOLARCHARGER BOARD WIRING |  | CHECK THE SOLAR CHARGER BOARD COMPONENTS DAMAGE | REPLACE THE SOLAR CHARGER BOARD
|  |  | BATTERY CHARGING VOLTAGE IS LOW | CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V | REPLACED THE CHARGER ADAPTER | REMOVE THE BATTERY AND CHECK THE VOLTAGE ADJUST TRIM.POT. OF THE SOLAR CHARGER BOARD | REPLACE THE CHARGER BOARD
|  |  | BATTERY CHARGING VOLTAGE IS TOO HIGH | REMOVE THE BATTERY AND CHECK THE VOLTAGE ADJUST TRIM.POT. OF THE SOLAR CHARGER BOARD | REPLACE THE CHARGER BOARD
|  |  | BATTERY CHARGING VOLTAGE IS CORRECT BUT BATTERY NOT CHARGING | CHECK THE BATTERY CONDITION | REPLACE THE BATTERY | CHECK THE SOLAR CHARGER BOARD COMPONENTS DAMAGE | REPLACE THE SOLAR CHARGER BOARD
|  |  | BATTERY CHARGING INDICATION LED NOT WORKING | CHECK THE BATTERY CONDITION | REPLACE THE BATTERY | CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY | REPLACE THE LED AND CABLE ASSY. | CHECK THE VOLTAGE OF THE CURREN INDICATION LINE | REPLACE THE SOLAR CHARGER BOARD
|  |  | LOW VOLTAGE CUTOFF NOT WORKING | CHECK THE LOW VOLTAGE CUTOFF RELAY(K2) IS WORKING(POWER ON SWITCH SUDDEN ON/OFF/ON) | REPLACE THE SOLARCHARGER BOARD
|  |  | LOW VOLTAGE CUTOFF LED NOT WORKING | CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY | REPLACE THE LED AND CABLE ASSY. | CHECK THE VOLTAGE OF THE CUTOFF INDICATION LINE | REPLACE THE SOLAR CHARGER BOARD
|  |  | MACHINE FULL TIME WORKING FROM BATTERY(BATTERY NOT FULL) | CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V | REPLACE THE CHARGER ADAPTER | CHECK THE SOLAR CHARGER BOARD LOW VOLTAGE SETTING TRIM POT. | REPLACE THE SOLARCHARGER BOARD | CHECK THE CHARGER ADAPTER DETECTION RELAY(K1) IS WORKING | REPLACE THE SOLARCHARGER BOARD
|  |  | MACHINE NOT WORKING FROM THE EXTERNAL CONNECTED BATTERY | CHECK THE EXTERNAL BATTERY VOLTAGE ABOVE 11.5V | CHARGE THE EXTERNAL BATTERY | CHECK THE RELAY(K1) IS WORKING (USE THE CHARGER ADAPTER) | REPLACE THE SOLAR CHARGER BOARD | CHECK THE LOW VOLTAGE CUTOFF TRIM POT. OF THE SOLAR CHARGER BOARD | REPLACE THE SOLAR CHARGER BOARD
|  | COMPACT ADAPTER | ADAPTER OUTPUT VOLTAGE IS ZERO (LSE V3/ECO V) | CHECK THE STATUS LED IS GLOWING | OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) | CHECK THE AC CORD | REPLACE AC CORD | CHECK THE ADAPTER PCB | REPLACE THE ADAPTER PCB
|  |  | ADAPTER OUTPUT VOLTAGE IS ZERO (LSE V3/ECO V) BUT LED IS GLOWING | CHECK THE OUTPUT DC CONNECTER OF THE ADAPTER | REPLACE THE OUTPUT DC CABLE
|  |  | ADAPTER OUPUT VOLTAGE SHOWING 8V TO 9V (LOW VOLTAGE) | OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) | REPLACE THE ADAPTER PCB | REPLACE THE ADAPTER PCB
|  | CHARGER ADAPTER | ADAPTER OUTPUT VOLTAGE IS ZERO (LSE-S V3/ECOD/ECOSV) | CHECK THE STATUS LED IS GLOWING | OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) | CHECK THE AC CORD | REPLACE AC CORD | CHECK THE CHARGER ADAPTER PCB | REPLACE THE CHARGER ADAPTER PCB
|  |  | ADAPTER OUTPUT VOLTAGE IS ZERO (LSE-S V3/ECOD/ECOSV) BUT LED IS GLOWING | CHECK THE OUTPUT DC CONNECTER OF THE ADAPTER | REPLACE THE OUTPUT DC CABLE
|  |  | ADAPTER OUPUT 12 VOLTAGE SHOWING 8V TO 9V (LOW VOLTAGE) AND 24V IS 24V | OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) | REPLACE THE CHARGER ADAPTER PCB | REPLACE THE CHARGER ADAPTER PCB
|  |  | ADAPTER OUPUT 12 VOLTAGE SHOWING 12V AND 24V IS LOW VOLTAGE/ZERO. | OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) | REPLACE THE CHARGER ADAPTER PCB | REPLACE THE CHARGER ADAPTER PCB
|  | ECOD-DPST BOARD | SYTEM ON TIME "PLEASE WAIT….." NOTIFICATION CONTINUIOSLY SHOWING | CHECK THE SD CARD IS OK | REPLACE THE SD CARD
|  |  | DISPLAY BLANK | CHECK THE DPST BOARD BREATHING LED IS BLINKING | REPLACE THE DPST BOARD | CHECK THE DIRECTION OF THE DISPLAY FRC CABLE CONNECTOR | CONNECT THE FRC CONNECTOR AT OPPOSITE DIRECTION | CHECK THE DISPLAY FRC CABLE QUALITY AND LCD QUALITY | REPLACE THE DISPLAY
|  |  | DISPLAY SHOWN HALF PORTION ONLY | CHECK THE DPST TO LCD FRC CABLE AND CONNECTOR QUALITY | REPLACE THE DPST TO LCD CABLE | CHECK THE DISPLAY MODULE FUNCTIONALITY | REPLACE THE LCD MODULE | CHECK THE DPST BOARD FUNCTIONALITY | REPLACE THE DPST BOARD
|  |  | RTC TIME NOT SAVED - RESTART TIME TIME CHANGED(RE-SETTED) | CHECK THE CMOS BATTERY VOLTAGE ABOVE 2V | REPLACE THE CMOS BATTERY | CHECK THE U5 AND PROCESSOR CONNECTING LINE SOLDERING | RE-SOLDER THAT CONNECTION
|  |  | RTC TIME NOT RUNNING | CHECK THE U5 IC CRYSTAL - Y3 AND ALSO CHECK C37 & C38 | REPLACE THE Y3 CRYSTAL AND C37&C38
|  |  | DPST KEY PAD NOT WORKING | CHECK THE DIRECTION OF THE KEYPAD FRC-10 CABLE CONNECTOR | CONNECT THE FRC-10 CONNECTOR AT OPPOSITE DIRECTION | CHECK THE DPST KEYPAD QUALITY | REPLACE THE DPST KEYPAD | CHECK THE DPST BOARD QUALITY | REPLACE THE DPST BOARD
|  |  | USB INITIALIZATION ERROR SHOWN | CHECK THE DPST TO BACKPANEL CABLE QUALITY AND DIRECTION OF THE FRC CONNECTOR | REPLACE THE FRC CABLE OR CONNECT THE FRC CONNECTOR AT OPPOSITE DIRECTION | CONNECT THE NEW BACK PANEL BOARD AND CHECK | REPLACE THE DPST BOARD
|  |  | WEIGHING SCALE NOT WORKING | CHECK THE REQUIRED SETTINGS IS APPLIED | SET THE REQUIRED SETTINGS VALUES | CHECK THE SCALE INTERFACING CABLE CONNECTIVITY | REPLACE THE CABLE | CHECK THE BACK PANEL TO DPST SCALE CABLE | REPLACE THE BACK PANEL TO DPST SCALE CABLE | CHECK THE BACKPANEL PCB FUNCTANALITY | REPLACE THE BACKPANEL PCB | CHECK THE DPST BOARD FUNCTANALITY | REPLACE THE DPST BOARD
|  |  | COMPUTER OUTPUT NOT PRESENT | CHECK THE DATAOUTPUT FORMAT OF THE ANALYZER SETTINGS | SET THE REQUIRED SETTINGS VALUES | CHECK THE MACHINE COMPUTER OUT TO COMPUTER CABLE CONNECTIVITY | REPLACE THE COMPUTER OUT TO COMPUTER CABLE | CHECK THE ANALYZER MAIN BOARD TO BACKPANEL INTERFACING CABLE | REPLACE THE ANALYZER MAIN BOARD TO BACKPANEL INTERFACING CABLE | CHECK THE BACKPANEL PCB FUNCTANALITY | REPLACE THE BACKPANEL PCB | CHECK THE ANALYZER MAIN BOARD FUNCTANALITY | REPLACE THE ANALYZER MAIN BOARD
|  |  | EXTERNAL DISPLAY DATA NOT PRESENT | CHECK THE MACHINE EXTERNAL DISPLAY OUT TO DISPLAY CABLE CONNECTIVITY | REPLACE THE MACHINE EXTERNAL DISPLAY OUT TO DISPLAY CABLE | CHECK THE DPST BOARD TO BACKPANEL EXTERNAL DISPLAY CABLE | REPLACE THE DPST BOARD TO BACKPANEL EXTERNAL DISPLAY CABLE | CHECK THE BACKPANEL PCB FUNCTANALITY | REPLACE THE BACKPANEL PCB | CHECK THE DPST BOARD FUNCTANALITY | REPLACE THE DPST BOARD
|  |  | RATE NOT TAKEN FROM THE CHART / NOT VIEW | CHECK THE SD CARD IS OK | REPLACE THE SD CARD | CHECK THE RATE CHART FILE IS SAVED IN THE SD CARD | CHECK THE SAVED FILE TYPE AND DATAS FORMAT IN CSV | SAVE TO THE CORRECT FORMAT AND TYPE ( CSV ) OF FILE | CHECK THE DPST BOARD FUNCTANALITY | REPLACE THE DPST BOARD
|  |  | FARMER DETAILS NOT SHOWN | CHECK THE SD CARD IS OK | REPLACE THE SD CARD | CHECK THE RATE CHART FILE IS SAVED IN THE SD CARD | CHECK THE SAVED FILE TYPE AND DATAS FORMAT IN CSV | SAVE TO THE CORRECT FORMAT AND TYPE ( CSV ) OF FILE | CHECK THE DPST BOARD FUNCTANALITY | REPLACE THE DPST BOARD
|  |  | ANALYZER CONNECTIVITY ERROR SHOWN | CHECK THE ANALYZER TO DPST BOARD INTERFACING CABLE AND CONNECTOR QUALITY | REPLACE THE ANALYZER TO DPST BOARD INTERFACING CABLE | CHECK THE ANALYZER MAIN BOARD FUNCTANALITY | REPLACE THE ANALYZER MAIN BOARD | CHECK THE ANALYZER DPST BOARD FUNCTANALITY | REPLACE THE DPST BOARD
|  |  | WIFI/GSM MODULE ERROR SHOWN | CHECK THE WIFI/GSM SETTINGS DETAILS | ENTER THE REQUIRED SETTINGS | CHECK THE DPST TO BACKPANEL FRC CABLE AND CONNECTION DIRECTION | REPLACE THE FRC CABLE OR RE CONNECT THE FRC CABLE AT CORRECT DIRECTION | CHECK THE WIFI/GSM MODULE BREATHING LED IS WORKING | REPLACE THE WIFI/GSM MODULE | CHECK THE ANTENNA SMA CABLE CONNECTION QUALITY AND ANTENNA CONNECTOR QUALITY | REPLACE THE ANTENNA AND SMA CABLE
|  |  | SMS NOT SEND TO THE FARMER | CHECK THE GSM/SIM MODULE CONNECTIVITY (INSERTED SIM DIRECTION ) AND CHECK THE SIM VALIDITY | CORRECT THE MODULE/SIM CONNECTIVITY AND EXTEND THE SIM VALIDITY | CHECK THE FARMER DETAILS ENTERED ARE CORRECT | PROPERLY ENTER THE FARMER DETAILS | CHECK THE ENTERED SMS ENABLE SETTINGS ARE CORRECT | ENTER THE CORRECT SETTINGS
|  |  | TESTED RESULT TO CLOUD UPDATION ERROR SHOWN | CHECK THE WIFI/GSM CONNECTIVITY | ENTER THE PROPER SETTINGS AND SET THE WIFI/GSM MODULE CONNECTIVITY | CHECK THE ENTERED THE CLOUD DETAILS ARE CORRECT | PROPER CLOUD ADDRESS AND DETAILS ARE ENTER | CHECK THE SAME MACHINE ID AND MODEL DETAILS ARE ENTERED IN THE CLOUD APPLICATION-(MACHINE REGISTRATION) | ENTER THAT MACHINE ID AND MODEL IN THE WEB APPLICATION
|  |  | CHART FILE NOT SHOWN / FARMER DETAILS NOT SAVED / COLLECTION NOT SAVED | CHECK THE SD CARD FILE FORMAT TYPE- ( USE FAT SYSTEM ) | FORMAT THE SD CARD AT FAT FILE SYSTEM USE THE COMPUTER | CHECK THE SD CARD IS DETECTED THE MACHINE ( READ /WRITABLE ) | FORMAT THE SD CARD USE THAT MACHINE
|  |  | SAMPLE TESTING TIME THE COUNT CONTINUOUSLY COUNTED ( ABOVE 50 SECOND ) | CHECK THE DPST BOARD TO ANALYZER MAIN BOARD INTERFACE CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE DPST BOARD TO ANALYZER MAIN BOARD | CHECK THE DPST BOARD FUNCTIONALITY | REPLACE THE DPST BOARD | CHECK THE ANALYZER MAIN BOARD FUNCTIONALITY | REPLACE THE ANALYZER MAIN BOARD
|  |  | DPST BOARD TO ANALYZER MAIN BOARD COMMUNICATION NOT PROPERLY WORKING | CHECK THE ANALYZER MAINBOARD MACHINE MODEL SELECTION IS "ECOD"
|  |  | PENDRIVE DETECTION FAILURE | CHECK THE PENDRIVE FUNCTIONALITY AND PENDRIVE FILE FORMAT IS FAT FILE SYSTEM | CHANGE THE PENDRIVE FILE FORMAT SET THE FAT FILE SYSTEM | CHECK THE BACKPANEL BOARD FUCTIONALITY | REPLACE THE BACKPANEL BOARD
|  |  | EXTERNAL KEYBOARD NOT DETECTED | CHECK THE DPST BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION | REPLACE THE DPST TO BACKPANEL FRC CABLE AND RE CONNECT THE CABLE AT PROPER DIRECTION | CHECK THE BACKPANEL BOARD FUCTIONALITY | REPLACE THE BACKPANEL BOARD | CHECK THE DPST BOARD FUNCTIONALITY | REPLACE THE DPST BOARD
|  |  | EXTERNAL DISPLAY NOT WORKING | CHECK THE DPST BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION | REPLACE THE DPST TO BACKPANEL FRC CABLE AND RE CONNECT THE CABLE AT PROPER DIRECTION | CHECK THE BACKPANEL BOARD FUCTIONALITY | REPLACE THE BACKPANEL BOARD | CHECK THE DPST BOARD FUNCTIONALITY | REPLACE THE DPST BOARD | CHECK THE EXTERNAL DISPLAY INTERFACING CABLE | REPLACE THE EXTERNAL DISPLAY INTERFACING CABLE
|  |  | PRINTER NOT WORKING | CHECK THE PRINTER POWER SUPPLY - 8 TO 9V | CHANGE THE POWER CABLE OR PRINTER SUPPLY BOARD | CHECK THE PRINTER FUNCTIONALLY WORKING - SELF TEST PRINTOUT | CHANGE THE PRINTER | CHECK THE PRINTER SETTINGS | CHANGE THE SETTINGS | CHECK THE COMMUNICATION CABLE | CHANGE THE CABLE
|  | PUMP | PUMP NOT WORKING | CHECK THE ANALYZER MAINBOARD BREATHING LIGHT IS WORKING | REPLACE THE ANALYZER BOARD | CHECK THE MOTOR CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE MOTOR | CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE CABLE | CHECK THE MOTOR MOTION SENSOR BOARD FUNCTANALITY | REPLACE THE MOTOR MOTION SENSOR BOARD | CHECK THE MOTOR DRIVER MODULE QUALITY | REPLACE THE MOTOR DRIVER MODULE | CHECK THE ANALYZER MAINBOARD FUNCTIONALITY | REPLACE THE ANALYZER MAINBOARD
|  |  | PUMP NUT AND BOLT MOVE TO ONESIDE ONLY AND HIT INTO THE CLAMP | CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE MOTOR MOTION SENSING BOARD CABLE | CHECK THE MOTOR MOTION SENSING BOARD FUNCTIONALITY | REPLACE THE MOTOR MOTION SENSING BOARD | CHECK THE ANALYZER MAINBOARD FUNCTIONALITY | REPLACE THE ANALYZER MAINBOARD
|  |  | PUMP SENSE ERROR | CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE MOTOR MOTION SENSING BOARD CABLE | CHECK THE MOTOR MOTION SENSING BOARD FUNCTIONALITY | REPLACE THE MOTOR MOTION SENSING BOARD | CHECK THE ANALYZER MAINBOARD FUNCTIONALITY | REPLACE THE ANALYZER MAINBOARD
|  | ECOD/ECOSV/LSES V3 BATTERY PROBLEMS | LOW BATTERY ERROR SHOWN | CHECK THE BATTERY VOLT ABOVE 11.5V | REPLACE THE BATTERY | CHECK THE SOLARCHARGER BOARD TO ANALYZER INTERFACE CABLE CONNECTIVITY | REPLACE THE SOLARCHARGER BOARD TO ANALYZER INTERFACE CABLE | CHECK THE ANALYZER MAINBOARD FUNCTIONALITY | REPLACE THE ANALYZER MAINBOARD | CHECK THE SOLAR CHARGER BOARD FUNCTANALITY | REPLACE THE SOLAR CHARGER BOARD
|  |  | BATTERY FULL LED INDICATION NOT WORKING | CHECK THE ANALYZER MAINBOARD IS COMPLETED THE WARMING PROCESS | APPLY TO CANCELLED THE WARMING PROCESS | CHECK THE VOLTAGE OF THE BATTERY ABOVE 14.3V | CHARGE THE BATTERY AT FULL | CHECK THE CHARGER BOARD TO ANALYZER MAINBOARD INTERFACING CABLE AND CONNECTOR QUALITY | REPLACE THE CHARGER BOARD TO ANALYZER MAINBOARD INTERFACING CABLE | CHECK THE VOLTAGE OF THE SOLAR CHARGER BOARD TO BATTERY CABLE LINE WITHOUT CONNECT THE BATTERY (14.45V) | ADJUST TRIMPOT OF THE SOLAR CHARGER BOARD AND SET AT THE CHARGING VOLTAGE (SET AT 14.45V). | CHECK THE BATTERY FULL INDICATION LED IS WORKING | REPLACE THE BATTERY FULL INDICATION LED | CHECK THE ANALYZER MAIN BOARD TO LED CABLE AND CONNECTOR CONNECTIVITY | REPLACE THE CABLE | CHECK THE ANALYZER MAINBOARD FUNCTIONALITY | REPLACE THE ANALYZER MAINBOARD
|  | ANALYZER MAINBOARD RELATED FAILURE - LSE V3,ECOV,ECOSV | T2/TEMP.SET ERROR | CHECK THE LEAKAGE/BLOCK OF THE SAMPLE SUCKING SECTIONS |  | CHECK THE INLET PIPE IS BROKEN OR BLOCKED | REPLACE THE INLET PIPE | CHECK THE INLET TUBE TO PREHEATER SILICON TUBE BROKEN OR LOOSE | REPLACE THE SILICON TUBE OR TIE THE SILICON TUBE | CHECK THE PRE HEATER TO SENSOR SILICON TUBE IS BROKEN OR LOOSE | REPLACE THE SILICON TUBE OR TIE THE SILICON TUBE | CHECK THE SENSOR TO SYRINGE SILICON TUBE BROKEN OR LOOSE | REPLACE THE SILICON TUBE OR TIE THE SILICON TUBE | CHECK THE SYRINGE IS DAMAGE | REPLACE THE SYRINGE | CHECK THE BELLOW TO L-PLUG SILICON TUBE IS BROKEN OR LOOSE AND ALSO CHECK ANY BUR OR OTHER PARTICLE HOLDED IN THE INNER SIDE OF THE SILICON TUBE AND BELLOW INLET SIDE | REPLACE THE SILICON TUBE OR BUR | CHECK THE L-PLUG BROKEN AND ALSO CHECK THE L-PLUG O-RING QUALITY | REPLACE THE L-PLUG OR O-RING | CHECK THE BELLOW QUALITY | REPLACE THE BELLOW | CHECK THE SENSOR INNER BOX SIDE ANY LEAKAGE/BLOCK RELATED ISSUES | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | BOX TEMP ERROR | CHECK THE JUMPER QUALITY OF THE MAINBOARD TO SENSOR BOARD SUPPLY LINE | REPLACE THE JUMPER | CHECK THE SENSOR BOARD FUNCTIONALITY | REPLACE THE SENSOR BOARD | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | SAMPLE NOT FOUND/AIR IN MILK | CHECK THE SAMPLE SUCKING SECTION ANY LEAKAGE/BLOCK | REPLACE THE DAMAGE ITEM | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | PLUNGE IN WATER/WATER IN SENSOR | CHECK THE SAMPLE LIQUID PRESENTS IN THE SENSOR TUBE | REMOVE THE SAMPLE LIQUID | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | HOT SAMPLE ERROR | CHECK THE TEST SAMPLE TEMPERATURE BELOW 40 DEGREE | REPLACE THE SAMPLE | CHECK THE SENSOR BOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | COUNT ZERO / RESULT ZERO | CHECK THE SENSOR BLOCK / LEAKAGE AND TEST COUNT VALUE | CLEAN THE SENSOR | CHECK THE SENSOR GAIN SET AT PROPER VALUE | SET THE NORMAL GAIN SETTINGS | CHECK THE SENSOR BOARD FUNCTIONALITY | REPLACE THE SENSOR BOARD | CHECK MAIN BOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | READING VARIATION | CHECK THE SENSOR AND SILICONE TUBES ARE CLEANED | CLEAN THE SENSOR AND SILICONE TUBES | CHECK THE SENSOR GAIN SET AT PROPER VALUE | SET THE NORMAL GAIN SETTINGS | CHECK THE PROPER OFFSET APPLY AND TEST READING | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR | CHECK MAIN BOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | FAT SHOWN IN WATER | CHECK THE ANALYZER OTHER MILK SAMPLES TEST READINGS ARE 'OK' | IF NOT OK CHECK THE CALIBRATION DETAILS | CHECK THE SENSOR AND TUBE WELL CLEANED | TO CLEAN THE SENSOR AND TUBE | APPLY WATER ZERO OPERATION AND TEST AGAIN | ENSURE THE PAIRED MAIN BOARD IS FUNCTIONALLY OK OTHERWISE BOARD REPLACE
|  |  | LCD CONTINIOUSLY SHOWING COMPUTER MODE | CHECK THE FRONTPANEL KEYPAD KEYS QUALITY | REPLACE THE KEYPAD | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | LCD CONTINIOUSLY SHOWING VERSION | CHECK THE FRONTPANEL KEYPAD KEYS QUALITY | REPLACE THE KEYPAD | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | WARMING STARTING TIME MACHINE WILL OFF/FUSE BURN | CHECK THE MAINBOARD FUNCTIONALITY - MOSFET FAILURE | REPLACE THE MAIN BOARD AND SENSOR BOARD PAIR | CHECK THE SENSOR BOARD COIL IS BURN | REPLACE THE MAIN BOARD AND SENSOR BOARD PAIR | CHECK THE ON/OFF SWITCH FUNCTIONALITY AND FUSE CARRER&FUSE IS TIGHT | REPLACE THE SWITCH OR RE-SET THE FUSE AND FUSE CARRIER | CHECK THE POWER CABLE OR CONNECTOR IS LOOSE | REPLACE THE FAULTY ITEM
|  |  | PENDRIVE AND KEYBOARD NOT DETECTED | CHECK THE MAINBOARD USB SETTING IS CORRECT | CORRECT THE USB SETTINGS | CHECK THE USB CONNECTOR AND INTERFACING CABLE | REPLACE THE CONNECTOR OR CABLE | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR BOARD PAIR
|  |  | WIFI/GSM MODULE ERROR SHOWN | CHECK THE WIFI/GSM SETTINGS DETAILS | ENTER THE REQUIRED SETTINGS | CHECK THE MAINBOARD TO MODULES OR BACKPANEL FRC CABLE AND CONNECTION DIRECTION | REPLACE THE CABLE OR RE CONNECT THE CABLE AT CORRECT DIRECTION | CHECK THE WIFI/GSM MODULE BREATHING LED IS WORKING | REPLACE THE WIFI/GSM MODULE | CHECK THE ANTENNA SMA CABLE CONNECTION QUALITY AND ANTENNA CONNECTOR QUALITY | REPLACE THE ANTENNA AND SMA CABLE
|  |  | FRONTPANEL KEYPAD NOT WORKING | CHECK THE KEYPAD,INTERFACING CABLE AND CONNECTOR QUALITY | REPLACE THE KEYPAD,INTERFACING CABLE AND CONNECTOR | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | LCD NOT WORKING - BLANK DISPLAY | CHECK THE LCD,INTERFACE CABLE AND CONNECTOR QUALITY | REPLACE THE LCD,INTERFACING CABLE AND CONNECTOR | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | RTC TIME NOT RUNNING | CHECK THE RTC MODULE IC CRYSTAL AND ALSO CHECK THE FILTER CAPACITORS | REPLACE THE CRYSTAL AND CAPACITORS | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | COMPUTER OUTPUT NOT PRESENT | CHECK THE DATAOUTPUT FORMAT OF THE ANALYZER SETTINGS | SET THE REQUIRED SETTINGS VALUES | CHECK THE MACHINE COMPUTER OUT TO COMPUTER CABLE CONNECTIVITY | REPLACE THE COMPUTER OUT TO COMPUTER CABLE | CHECK THE MAINBOARD FUNCTIONALITY | REPLACE THE MAINBOARD AND SENSOR PAIR
|  |  | RTC TIME NOT SAVED - RESTART TIME TIME CHANGED(RE-SETTED) | CHECK THE CMOS BATTERY VOLTAGE ABOVE 2V | REPLACE THE CMOS BATTERY | CHECK THE RTC MODULE AND MAINBOARD FUNCTIONALITY | REPLACE THE RTC MODULE OR MAINBOARD PAIR
|  |  | SMS NOT SEND TO THE FARMER | CHECK THE GSM/SIM MODULE CONNECTIVITY (INSERTED SIM DIRECTION ) AND CHECK THE SIM VALIDITY | CORRECT THE MODULE/SIM CONNECTIVITY AND EXTEND THE SIM VALIDITY | CHECK THE FARMER DETAILS ENTERED ARE CORRECT | PROPERLY ENTER THE FARMER DETAILS | CHECK THE ENTERED SMS ENABLE SETTINGS ARE CORRECT | ENTER THE CORRECT SETTINGS
|  |  | TESTED RESULT TO CLOUD UPDATION ERROR SHOWN | CHECK THE WIFI/GSM CONNECTIVITY | ENTER THE PROPER SETTINGS AND SET THE WIFI/GSM MODULE CONNECTIVITY | CHECK THE ENTERED THE CLOUD DETAILS ARE CORRECT | PROPER CLOUD ADDRESS AND DETAILS ARE ENTER | CHECK THE SAME MACHINE ID AND MODEL DETAILS ARE ENTERED IN THE CLOUD APPLICATION-(MACHINE REGISTRATION) | ENTER THAT MACHINE ID AND MODEL IN THE WEB APPLICATION
|  |  | RATE NOT TAKEN FROM THE CHART / NOT VIEW | CHECK THE RATE CHART SETTINGS ENTERED PROPERLY | ENTER THE CORRECT SETTINGS
|  |  | FARMER DETAILS NOT SHOWN | CHECK THE FARMER DETAILS ARE ENTERED CORRECTLY | ENTER THE CORRECT DETAILS
|  |  | WEIGHING SCALE NOT WORKING | CHECK THE REQUIRED SETTINGS IS APPLIED | SET THE REQUIRED SETTINGS VALUES | CHECK THE MAIN BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION | REPLACE THE MAIN TO BACKPANEL FRC CABLE AND RE CONNECT THE CABLE AT PROPER DIRECTION | CHECK THE BACKPANEL PCB FUNCTANALITY | REPLACE THE BACKPANEL PCB | CHECK THE MAIN BOARD FUNCTANALITY | REPLACE THE MAINBOARD AND SENSOR PAIR | CHECK THE WEIGHING SCALE INTERFACING CABLE | REPLACE THE WEIGHING SCALE INTERFACING CABLE
|  |  | EXTERNAL DISPLAY NOT WORKING | CHECK THE MAIN BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION | REPLACE THE MAIN TO BACKPANEL FRC CABLE AND RE CONNECT THE CABLE AT PROPER DIRECTION | CHECK THE BACKPANEL BOARD FUCTIONALITY | REPLACE THE BACKPANEL BOARD | CHECK THE MAIN BOARD FUNCTIONALITY | REPLACE THE MAIN BOARD | CHECK THE EXTERNAL DISPLAY INTERFACING CABLE | REPLACE THE EXTERNAL DISPLAY INTERFACING CABLE
|  | OTHERS | ANALYZER NOT WORKING | CHECK THE FUSE | REPLACE THE FUSE | CHECK THE ADAPTER / BATTERY VOLTAGE ABOVE 11.5V | REPLACE THE ADAPTER / BATTERY | CHECK THE ON/OFF SWITCH FUNCTIONALITY | REPLACE THE ON/OFF SWITCH | CHECK THE INTERNAL BARDS SUPPLY LINE IS SHORTCIRCUIT | REPLACE THE SHORTED BOARD | CHECK THE INTERNAL SUPPLY LINE OPEN/LOOSE FAULT | REPLACE THE FAULTY CONNECTOR/CABLES

## WATER ZERO
| SI.NO. | STEPS
WATER ZERO | 1 | First to clean the machine use the manual plunging method.(Use to 60 Degree hot water)
| 2 | To clean the machine ( Auto weekly cleaning Process ), use the weekly solution.
| 3 | To test a sample milk at one times.( FAT- 2 TO FAT- 4 Nearest value)
| 4 | To clean the machine use the manual plunging method.
| 5 | To apply the water zero setting process,( To use only fresh and normal temerature water.)

## EASY CORRECTION
|  | SI.NO. | PROCESS STEPS | STATUS | ACTION
STEP-1 | MACHINE CLEAN AND WATER ZERO | 1 | First to clean the machine use the manual plunging method.(Use to 60 Degree hot water)
|  | 2 | To clean the machine( Auto weekly cleaning Process ), use the weekly solution.
|  | 3 | To test a sample milk at one times. ( FAT- 2 TO FAT- 4 Nearest value)
|  | 4 | To clean the machine use the manual plunging method.
|  | 5 | To apply the water zero setting process,( To use only fresh and normal temerature water.)
PROCESS STEPS |  |  |  | RESULT | ACTION
STEP-2 | EASSY CORRECTION SETTING | 1 | Take 3-types of samples (COW channel - F 2.00 , F 4.00 & F 6.5 - nearest values are acceptable.)
|  | 2 | First F 2.00 sample circulate - ( Lower value milk use to first circulation )
|  | 3 | Then test whole samples and find out the variation point ( Sample).
|  | 4 | Then test the deviated sample again and follow the below steps. (Examples shown below)
|  | 5 | Test Result-1-Then test Each sample and note the result ( eg: result F 2.00 shown 1.8 / F 4.00 shown 4 / F 6.5 shown 6.5 ) | Deviation only at the lower side - F 2.00 (F1.8) | FIRST TEST THAT DEVIATED SAMPLE ( F1.80 ) THEN APPLY THE EASSY CORRECTION-F+0.2
|  | 6 | Other Test Result-2-Then test Each sample and note the result ( eg: result F 2.00 shown 2.00 / F 4.00 shown 3.8 / F 6.5 shown 6.5 ) | Deviation only at the - F 4.00 ( F3.8 ) | FIRST TEST THAT DEVIATED SAMPLE ( F 3.75 ) THEN APPLY THE EASSY CORRECTION-F+0.25
|  | 7 | Again test the above mensioned 3-types sample and check the Result ( COW channel - F 2.00 , F 4.00 & F 6.5 )
|  | 8 | If any changes are identified, repeat the above process.
*** Same Process follows for applying the corrections at the BUFFALO and MIX channels ***
*** If the common offset is present, First remove that common offset then apply the eassy correction
STEP-3 | MACHINE CLEAN AND WATER ZERO | 1 | First to clean the machine use the manual plunging method.(Use to 60 Degree hot water)
|  | 2 | To clean the machine( Auto weekly cleaning Process ), use the weekly solution.
|  | 3 | To test a sample milk at one times. ( FAT- 2 TO FAT- 4 Nearest value)
|  | 4 | To clean the machine use the manual plunging method.
|  | 5 | To apply the water zero setting process,( To use only fresh and normal temerature water.)

## CORRECTION
|  | SI.NO. | PROCESS STEPS
STEP-1 | MACHINE CLEAN AND WATER ZERO | 1 | First to clean the machine use the manual plunging method.(Use to 60 Degree hot water)
|  | 2 | To clean the machine( Auto weekly cleaning Process ), use the weekly solution.
|  | 3 | To test a sample milk at one times.( FAT- 2 TO FAT- 4 Nearest value)
|  | 4 | To clean the machine use the manual plunging method.
|  | 5 | To apply the water zero setting process,( To use only fresh and normal temerature water.)
STEP-2 | OFFSET SETTING | 1 | To test a sample milk (Reference Milk) at one times. (Sample same ambient temperature)
|  | 2 | To test Reference Milk second sample.. (Sample same ambient temperature) and apply the offset value.
|  | 3 | To test same Reference Milk third sample. (Sample same ambient temperature) and ensure the result is correct.
|  | 4 | If the test results are significantly deviated or incorrect, restore the machine (channel) settings and repeat the offset procedure.
5
6
STEP-3 | MACHINE CLEAN AND WATER ZERO | 1 | First to clean the machine use the manual plunging method.(Use to 60 Degree hot water)
|  | 2 | To clean the machine( Auto weekly cleaning Process ), use the weekly solution.
|  | 3 | To test a sample milk at one times.( FAT- 2 TO FAT- 4 Nearest value)
|  | 4 | To clean the machine use the manual plunging method.
|  | 5 | To apply the water zero setting process,( To use only fresh and normal temerature water.)

## ECOD SETTINGS
NO. | MENU - PAGE | NO. | MENU LIST | NO. | SUB MENU | NEXT=>
1 | HOME PAGE | 1 | FARMER ID AND NAME | 1 | PROTEIN,LACTOSE,SALT AND TEMP.
|  | 2 | BATTERY LEVEL INDICATION
|  | 3 | COLLECTION SHIFT NAME-MOR/EVEN
|  | 4 | NUMPER OF TEST FINISHE IN A SHIFT COUNT
|  | 5 | TEST PARAMETERS-FAT,SNF,CLR……ETC
|  | 6 | RATE,WEIGHT AND TOTAL AMOUNT
RIGHT ARROW KEY ACTION MENU =>
|  | 7 | POP UP INDICATIONS :
KEY BOARD DETECTION,PENDRIVE DETECTETION AND WARNING INDICATION-AUTO CLEANING,LOW BATTERY ETC
2 | MENU PAGE-1 | 1 | DPU MENU | 1 | VIEW REPORT | COLLECTION REPORT | CUSTOMER REPORT | SALES REPORT | BILL SUMMARY | DESPATCH REPORT
|  |  |  | 2 | SELECT RATE CHART - CHART CREATION | COW,BUFF OR MIX CHART SELECTION | RATE SELECT FROM-FAT ONLY,CLR ONLY,FAT+CLR AND FAT+SNF | CHART GENERATION | CLR EQUATION | AUTO CHART - RATE SELECTION
|  |  |  | 3 | VIEW RATE CHART | COW,BUFF OR MIX CHART VIEW
|  |  |  | 4 | WEIGHNING SCALE SETTINGS | SCALE TYPES | MANUAL SETTINGS | TARE ACTION - BEFORE/AFTER RESULT | AUTO TARE ENABLE/DISABLE | WEIGHT CONVERTION-KG TO LITER | WEIGHNING SCALE LIVE WEIGHT SELECTION
|  |  |  | 5 | FILE MANAGEMENT | COPY FILE-MACHINE TO PENDRIVE/PENDRIVE TO MACHINE | DELETE FILES FROM MACHINE | FILE UPDATE FROM SERVER
|  |  |  | 6 | DECIMAL POINT SELECTION | 1 OR 2 DECIMAL POINT SELECTION
|  |  |  | 7 | SETTINGS-SYSTEM | FARMER DETAILS ENTRY | SYSTEM SETTINGS ENTRY | BONUS/RATE CHANGE SETTING | AUTO SMS/CLOUD UPDATE ENABLE/DISABLE | PENDING SMS REMOVE SELECTION | SMS ALERT ACTIVATION ENTRY | GPRS/WIFI SETTINGS | IN/OUT AND DISPLAY ACTIVATION ENTRY
|  |  |  | 8 | SHIFT TIME SETTINGS | COLLECTION SHIFT CHANGE ENTRY | SHIFT BONUS ENTRY
|  | 2 | ANALYZER SETTINGS | 1 | ANALYZER SETTINGS - CHANEL SETTINGS | CHANNEL SETTINGS | SYSTEM SETTINGS ENTRY | BACKUP/COPY
|  |  |  | 2 | LAST TEST RESULT | SHOWN LAST TEST RESULT
|  |  |  | 3 | TEST/CLEANING/STATUS LOG DETAILS | SHOWN NUMBER OF TEST / CLEANING DONE
|  |  |  | 4 | SENSOR TEST | SENSOR BOARD FUNCTIONAL TEST
|  | 3 | ID_WEIGHT ENTRY | 1 | FAST COLLECTION ENTERING WINDOW | FIRST FARMER ID AND SAMPLE WEIGHT ENTER AND THE SAMPLE TEST PARAMETERS ARE SAVED AT THE NEXT FREE TIME
|  | 4 | FACTORY SETUP | 1 | OFFICE PURPOSE
|  | 5 | MACHINE SETTINGS DETAILS | 1 | OFFICE PURPOSE
|  | 6 | MODULE TEST-GSM/WIFI | 1 | TROUBLESHOOTING | GSM MODULE MESSAGE READ AND DELTE PURPOSE | 1.MESSAGE READ,2.PHONE CALLING**,3.SERVER TEST** (IN FUTURE **)
|  |  | DPU MENU=>SETTINGS | 1 | FARMER DETAILS
|  |  |  | 2 | SYSTEM SETTINGS | MACHINE ID ENTRY | LOCATION DETAILS ENTRY-DISTRICT,VILLAGE…ETC | PASSWORD CHANGE | DATE AND TIME ENTRY | MANUAL TESTING ENABLE/DISABLE | AUTO PRINTOUT AND PRINTOUT CONTENT ENTRY
|  |  |  | 3 | BONUS/RATE CHANGE ADDING | COMMON BONUS ADDING AND COMMON RATE CHANGE IN THE RATE CHART +/- DEVIATION
|  |  |  | 4 | AUTOMATIC SMS SEND/GPRS UPDATE EN/DISBL | AUTO MATIC PENDING SMS SEND ENABLE | AUTOMATICALLY UPDATE PENDING COLLECTION,CORRECTION,CHART AND PASSWORD | HTTP BACK- UPLOADING COLLECTION BACKUP ENABLE OR DISABLE WHEN THE CONNECTION BREAK TIME.
|  |  |  | 5 | REMOVE PENDING SMS | PENDING SAVED SMS CLEARING
|  |  |  | 6 | SMS ALERT ENABLE/DISABLE | TOTAL SMS ALERT DISABLE OR ENABLE | GROUP SMS ALERT-SUMMARY SMS SEND TO FIRST 1 TO 9 SAVED PERSONS (FARMER ENTRY LOCATION 1 TO 9 -ID'S)
|  |  |  | 7 | GPRS/WIFI SETTINGS | WIFI OR GSM SELECTION | CLOUD URL ENTRY (IN GPRS SETTINGS MENU) AND CONNECTION TOKEN ENTRY | WIFI SETTINGS - SSID AND PASSWORD ENTRY
|  |  |  | 8 | IN/OUT/DISPLAY SETTINGS | DISPLAY PARAMETERS-LCD SHOWING PARAMETERS ENABLE/DISABLE | IN/OUT- EXTERNAL DISPLAY FORMAT,ID BASED DOUBLE ENTRY ENABLE/DISABLE AND DESPATCH ID ENTRY | ROUD OFF ENABLE / DISABLE | CURRENCY SELECTION | BATTERY WARNING ENABLE/DISABLE | ECOD MODEL SELECTION V3 OR V4

## LSE V3 MENU
DATE | SOFT.VERSION | BUILD | TINY | PURPOSE | SELECTED MODE/SETTINGS | STEPS/MENU | IDENTIFICATION IN THE MENU/STEPS | ACTION/DETAILS | REMARK
2024-12-09 | 3.3.6 | 3 | 12 | IMPORTANT POINTS
|  |  |  | 1) EEPROM SELECT - INTERNAL FLASH |  |  | FACTORY SETUP
|  |  |  | 2) DOUBLE ENTRY - ENABLE |  |  | OTHER SETTINGS=>FARMER SETTINGS=>OK/CNCL KEY
|  |  |  | 3) FIRST CLEAR THE COLLECTION AND SALES |  |  | FRONT MENU
|  |  |  | 4) AFTER CALIBRATION IT TO BACKUP-INDIVIDUAL AND ALL CHANNELS |  |  | OTHER SETTINGS=>BACK/COPY/REST=>UP/DWN KEY
|  |  |  | 5) SET THE PRINTER SETTINGS - LEADING SPACE = 005 AND FEED SPACE = 005 |  |  | OTHER SETTINGS=>PRINTER SETTINGS=>UP/DWN KEY
|  |  |  | 6) DELETE ALL FARMER |  |  | OTHER SETTINGS=>FARMER SETTINGS=>UP/DWN KEY
|  |  |  | 7) EDIT DISPATCH ID |  |  | OTHER SETTINGS=>FARMER SETTINGS=>UP/DWN KEY | eg: 001001
|  |  |  | 8) SET THE CODE/OTP SELECTION AT "PASSCODE" |  |  | OTHER SETTINGS=>CODE/OTP=>UP/DWN KEY | OTP SELECTION ACTIVATE ONLY AFTER THE BLUETOOTH CONNECTION FUNCTIONALLY PERFECT
|  |  | 1) | RESULT MANUAL EDITABLE - ENABLE/DISABLE | REMOTE MODE | * | STAR SYMBOL SHOWN IN TEST MODE = ( Remote* ) |  | ITS WORKING ONLY REMOTE MODE
[7 empty] | AFTER RESULT COMING A SHORT BEEP COMING- EDIT IS ACTIVE-TO USE KEYBOARD AND EDIT IT | PRESS CANCEL KEY- SAMPLE WILL OUT
PRESS OK KEY ITS WILL SAVED
[5 empty] | TO ACTIVATE | STEP-1 | GO TO "ANALYZER SETTINGS"
[6 empty] | STEP-2 | EDIT CHANNEL
[6 empty] | STEP-3 | SELECT CORRECTION
[6 empty] | STEP-4 | PRESS OK/CANCEL KEY => "EDIT ENABLE" - SET AT "ENABLE" | NOT CONTROL INDIVIDUALLY -CHANNEL
ANY CHANNEL IS SET IT WILL WORKED
|  |  | 2) | SAVED RESULT EDITABLE | REMOTE MODE | STEP-1 | FIRST ACTIVETE THE "EDIT ENABLE" - SET AT "ENABLE"===ANALYZER SETTINGS=>EDIT CHANNEL=>ENTER CORRECTION=>EDIT ENABLE |  | ITS WORKING ONLY REMOTE MODE
[6 empty] | STEP-2 | ENTER "VIEW RESULT"
USE=UP/DOWN KEY THEN SELECT "EDIT" THEN ENTER "OK"
|  |  | 3) | DOUBLE ENTRY ENABLE | REMOTE MODE | STEP-1 | GO TO "OTHER SETTINGS" |  | ITS WORKING ONLY REMOTE MODE
[6 empty] | STEP-2 | ENTER "FARMER SETTINGS"
[6 empty] | STEP-3 | PRESS OK/CANCEL KEY => "DOUBLE ENTRY" - SET AT "ENABLE" |  | ITS ENABLE=ACCEPTED MORE TESTS AT THE SAME FARMER ID
[7 empty] | DISABLE |  | ITS DISABLE=ACCEPTED ONLY ONE TEST IN A SHIFT - SAME FARMER ID
|  |  | 4) | WEIGHING SCALE | REMOTE MODE | SETTINGS | 1) "WEIGHT IN COLLECTION" WEIGHT TAKEN IN THE TIME OF TEST-BEFORE/AFTER | WEIGHT VALUE TAKEN BEFORE OR AFTER THE SAMPLE TEST | ITS WORKING ONLY REMOTE MODE
[7 empty] | 2) "WEIGHT DETECT FROM" - FROM SCALE/ MANUAL | WEIGHT TAKEN FROM WEIGHN SCALE OR MANUAL ENTER
[7 empty] | 3) KG TO LITRE | KILOGRAM DATA TO LITRE CONVERTION ENEBLE/DISABLE
[7 empty] | 4) TARE - NO TARE,BEFORE AND AFTER COLLECTION | AUTOMATIC SCALE TARE ACTION
5) BAUD RATE & DATATYPE
[7 empty] | 6) READ DELAY | SCALE DATA SENDING TIME ITERVEL = SET AT NORMALLY 0.50 SECOND
( 5) "BAUD RATE & DATATYPE"
[7 empty] | 1) SERIAL "BAUD RATE" AUTO DETECTION/MANUAL SELECTION | 1) AUTO SELECTION = TO ENTER THE REFERENCE WEIGHT DATA FROM THE SCALE SHOWING DISPLAY | EG: THE SCALE DISPLAY SHOWN 123.45 LTR. THIS VALUE ENTER IN THAT SETTING
AFTER ENTER THE VALUE THE MACHINE AUTOMATICALLY SEARCHING THE SERIAL DATA STRUCTURE
AFTER THE DETECTION THE RAW DATA SHOWN IN THE LCD
[8 empty] | 2) MANUAL SELECTION - TOTAL 10-TYPE SELECTABLE MODELS AND 1-NOS (11 TH) MANUAL SETTABLE MODEL | DOWN KEY USE TO SELECT THE " BAUD RATE" AND UP KEY USED TO SELECT THE "TYPES OF MODELS"
****TO SELECT THE 11 TH MANUAL ENTERING MODEL*****
[8 empty] | 1) FIRST ENTER THE START CHARACTER CODE | THE CHARACTER CODE TABLE SHOWN AT THE NEXT SHEET ( ASCII CODE ) = EG: SCALE SERIAL PORT DATA IS "L00123.45 0D 0A"
"L" CHARACTER CODE IS = 076
TO ENTER THE 076 DATA INTO THE START CHARCOD AREA
[8 empty] | 2) NEXT ENTER THE " END CHARCODE " | " 0D " IS THE END CODE = THE ENTERED CODE VALUE IS " 013 "
[8 empty] | 3) NEXT ENTER THE " CHAR COUNT " | EG: "L00123.45 0D 0A" THE CHARECTERS ARE THE "00123.45" TOTAL COUNT = 7
[8 empty] | 4) NEXT IS THE DECIMAL POINT COUNT | EG: "L00123.45 0D 0A" THE DECIMAL POINT VALUES ARE "45" COUNT=2
|  |  | 5) | CHART SETTINGS | NORMAL/REMOTEMODE |  | ANALYZER SETTINGS=>EDIT CHANNEL=>RATE.CH.SETTINGS
TYPES CHART - METHOD AND IMPORT FILE
1) METHODS - 1,2,3,4 AND 5 ( AND EXTRA SLOTTED METHOD)
[5 empty] | RATE CHART SELECTION SETTINGS | STEP-1 | 1) RATE CHART ENABLE - ENABLE
[6 empty] | STEP-2 | 2) AUTO CHART SELECTION AND LIMIT SET - ( PURPOSE COW CHART TO BUFFALO CHART SEARCHING ) | EG: FAT LIMIT FAT = 7.0 | BELOW FAT 7.0 RATE TAKEN FROM COW CHART AND ABOVE 7.0 FATE THE RATE TAKEN FROM BUFFALO CHART
[6 empty] | STEP-3 | 3) HIGH-LOW FX RATE - " ENABLE " | A FIXED RATE ADDED IN THE RESULT FOR THE BELOW FAT LIMIT (eg: F 2.00 ) AND ABOVE FAT LIMIT (eg: F 12.00 )
[6 empty] | STEP-4 | 4) LOW LIMIT FAT - eg: F 2.0 | BELOW FAT 2.0 - A FIXED RATE TAKEN
[6 empty] | STEP-5 | 5) HIGH LIMIT FAT - eg: F 12.0
[6 empty] | STEP-6 | 6) ENTER LOW FAT FIXED RATE = ENTER THE RUPEES
[6 empty] | STEP-7 | 7) ENTER HIGH FAT FIXED RATE = ENTER THE RUPEES
[6 empty] | STEP-8 | 8) RATE CHART SELECTION - METHOD/IMPORT FROM PENDRIVE
[5 empty] | METHOD SELECTED |  | RATE TAKEN COMBINATION-FAT&SNF,FAT&CLR,FAT ONLY,CLR ONLY | ITS USE THE RATE TAKEN FROM THE FAT AND SNF READING BASIS
[5 empty] | IMPORT SELECT | STEP-1 | RATE TAKEN COMBINATION-FAT&SNF,FAT&CLR,FAT ONLY,CLR ONLY
[6 empty] | STEP-2 | 1) ENTER START FAT - FOR THE IMPORT FILE |  | IF ANY WRONG RATE SHOWN AT THE "CHECK RATE" OPERATION --- CHECK THE "AUTO CHART IS ENABLE" AND ALSO CHECK THAT RANGE
[6 empty] | STEP-3 | 2) ENTER END FAT - FOR THE IMPORT FILE
[6 empty] | STEP-4 | 3) ENTER START SNF - FOR THE IMPORT FILE
[6 empty] | STEP-5 | 4) ENTER END SNF - FOR THE IMPORT FILE
THEN IMPORT FILE FROM USB
ADDITIONAL FEATURES
|  |  | 1 | SALES |  | STEP-1 | SELECT THE "SALES" OPTION
[6 empty] | STEP-2 | 1) SELECT THE TYPE-COW,BUF OR MIX
[6 empty] | STEP-3 | 2) ENTER LITRE
[6 empty] | STEP-4 | 3) ENTER THE RATE | THE RATE IS BUFFERED AT LAST ENTERED VALUE
|  |  | 2 | CLEAR SAVE RESULT-COLLECTION AND SALES |  |  | INDIVIDUALLY COLLECTION AND SALES CLEARING OPTION
|  |  | 3 | SET CURRENCY |  |  | eg: INR,EUR,USD…ETC | SHOWN ONLY PRINTOUT
|  |  | 4 | TEST MODE: 1 | NORMAL MODE
|  |  |  | 2 | REMOTE MODE
|  |  |  | 3 | MANUAL COLLECTION | STEP-1 | 1) SELECT MANUAL TEST MODE
[6 empty] | STEP-2 | 2) SELECT CHANNEL
[6 empty] | STEP-3 | 3) ENTER THE ID,WEIGHT,FAT…..etc THEN CONTINUE "OK" PRESS (IT SAVED ONLY PRESS THE "OK" KEY)
|  |  |  | 4 | ID&WEIGHT ENTRY | STEP-1 | 1) SELECT THE ID&WEIGHT ENTRY TEST MODE-UP/KEY USE | FIRST CONFORM WHICH MODE=LIKE REMOTE SAMPLE TEST/MANUAL ENTRY | THEN SELECT THE ID&WEIGHT ENTRY
[6 empty] | STEP-2 | 2) ENTER THE ID DATA AND WEIGHT DATA
[6 empty] | STEP-3 | 3) AFTER SAVED THE ID&WEIGHT THEN SELECT "MILK TEST-SAVEDWEIGHT" MENU
[6 empty] | STEP-4 | 4) THEN ENTER THE SAMPLE DATAS ( FAT/SNF..ETC) -- MANUAL/REMOTE TEST
[6 empty] | STEP-5 | 5) AFTER ENTER THE SAMPLE READING THEN SELECT&ENTER THE " CANCEL-SAVEDWEIGHT" MENU
|  |  | 5 | SHIFT BONUS | REMOTE MODE | STEP-1 | GO TO "OTHER SETTINGS"
[6 empty] | STEP-2 | ENTER "FARMER SETTINGS"
[6 empty] | STEP-3 | USE=UP/DOWN KEY THEN SELECT "ENABLE" THEN ENTER THE VALUES THEN "OK"
|  |  | 6 | CLOUD CONNECTIVITY-SETTINGS | CLOUD SETTING | OTHER SETTINGS | 1) SELECT OTHER SETTINGS
2) SELECT CLOUD SETTINGS- ENTER-SHOWS THE MODEL NAME
3) ENTER THE CLOUD URL DETAILS
4) ENTER THE UPLOAD TOKEN DETAILS
5) ENTER THE CLOUD APPLICATION MACHINE REGISTERED NUMBER DETAILS
[5 empty] | SETTINGS CLOUD UPDATES | HOME PAGE | 1) UPDATE FROM CLOUD-CHARTS,CORRECTION,PASSWORD AND STATISTICS
[5 empty] | FARMER DETAILS UPDATE | OTHER SETTINGS | 1) SELECT FARMER SETTINGS
2) SELECT UPDATE FROM CLOUD ( DOWNLOAD FARMER DETAILS)
*** THE 336 SOFTWARE SUPPORT THE CHART DOWNLOADING IN THE FILE FORM

## Sheet5
FAT | Fats are made from individual fatty acid molecules attached to glycerol, a 3-carbon backbone. The most common type of fat is called a triglyceride, or triacylglycerol, which contains 3 fatty acids attached to the backbone and resembles a fork without the handle.
CLR | The corrected value or corrected lactometer reading (CLR) is obtained from the standard table for corresponding temperature.A lactometer is a tiny glass device used to measure the cleanliness or quality of milk. It determines the density of milk in relation to water. If the specific gravity of a milk specimen falls inside the acceptable parameters, the milk is considered pure.
SNF | SNF or solids not fat is that nutrient portion present in milk which is other than milk fat and water. It consists of protein (primarily casein and lactalbumin), carbohydrates (primarily lactose), and minerals (including calcium and phosphorus. When SNF is combined with milk fat, then it is called total solids.